## FMEA headings
Function | Failure mode | Effects | S | Cause(s) | O | Current controls | D | CRIT | RPN | Recommended actions | Responsibility | Action taken
vertical gravity compensation | spring yields | force at given location reduces | 2 | stretching spring past elastic limit | 1 | microcontroller varifies idle position | 1 | detection requires calibration of device for idle position and upper and lower limit | 2 | none
| spring breaks | no verticle gravity compensation will exist | 5 | fatigue of spring | 3 | none | 10 | N/A | 150 | Ensure spring selection and sizing considers fatigue requirements
| spring becomes misaligned | force at given location varies | 2 | spring is knocked | 5 | microcontroller varifies idle position | 1 | detection requires calibration of device for idle position and upper and lower limit | 10 | none
|  |  |  | design extends out of desired range | 1 | radial link prevents undesirable motion | 1 | detection requires calibration of device for idle position and upper and lower limit | 2 | none
|  | no verticle gravity compensation will exist | 5 | design extends out of desired range | 1 | user notices misalignment | 8 | difficult to dect, therefore must prevent | 40 | ensure guides are positioned to ensure spring cannot fall off device
stop device rotating about joint 1 | device rotates about joint 1 | device rotates undesirably about joint 1 - user must apply addition force during operation | 2 | device is positioned so that torque to resist rotation is less than torque applied by device | 6 | microcontroller varifies tool position is within desired range | 3 | desired tool position must be specified to reduce the input force from user required while still providing the necessary operating envelop | 36 | torque produced by counterweight or motor must be optimized to minimize potential to exceed range
|  | device rotates beyond implemented envelop - interaction with patient and objects not considered in design | 8 | device is positioned so that torque to resist rotation is less than torque applied by device | 6 | physical stopper positioned to limit rotation about joint 1 | 1 | all possible orientations of the device at the specified angles must be considered to ensure the tool cannot come into contact with objects outside operating vacinity | 48 | optimize physical stopper position